## Slide 1: MEG II実験 陽電子スペクトロメータの 3次元磁場測定
家城 佳、他MEG IIコラボレーション
日本物理学会 2018年秋季大会

## Slide 2: MEG II experiment
We search for cLFV decay,
Aimed branching ratio sensitivity: (90% C.L.) ⇔ Prediction from bSM (e.g. SUSY-seesaw):
Discovery of = Discovery of new physics!
μ+
Key concepts of the experiment:
- High intensity beam (/s) @ PSI
- High resolution detectors to distinguish signal from accidental BG
MeV
back-to-back
same timing
signal

## Slide 3: e+ spectrometer and COBRA magnet
Drift chamber
Timing counter
e+
+ COBRA magnet
e+ is swiped away quickly  hit rate reduces
(reused from MEG)
e+ spectrometer =
COnstant Bending RAdius magnet
- Superconducting magnet, ~1.2 T at center
- Gradient field
Same radius at different angles for same pe+
~2m

## Slide 4: Requirement for B-field measurement
Example of signal positron track
Bias in pZ vs. shift of B-map in Z
Default
Shift B-field map by Z +10cm
Resolution will improve in MEG II:
 B-field uncertainty should be <~0.1%
Alignment of the sensor is important
 e.g. If B-field map is shifted in Z by 0.5 mm, pz is biased by ~40 keV, vertex Z is biased ~0.2 mm

## Slide 5: Difficulties in gradient B-field measurement
Sensor position must be known well
Some commercial sensors are too large
e.g. NMR probe is usually too big
Position shift will make a bias in e+ reconstruction.
Gradient B-field  Strength and direction of the B-field changes at different place.
Bz
sensor Z
θ
Measure the field in 3 directions.  Angle misalignment makes a large effect
sensor X
0  Bz
large effect
Bz  Bz
small effect

## Slide 6: Previous measurements (for MEG)
A. Commercial 3D Hall sensor + moving wagon Angle error was large (9 mrad)  Only BZ was used. BR and BΦ were calculated from Maxwell equations.
B. 1D Hall sensor + rotating stage + moving arms  Cancel out the effect of misalignment by rotating sensor.
 Data taking time was long (~1 week).
It was also necessary to consider the
misalignment of the rotating stage
B
B
-B
Data points (XY view)
Hall sensor
+ rotators
Difference between measured and calculated field was ~0.2% (in RMS)

## Slide 7: Improvements in new measurement
4 mm
Hall sensor x6
“Hall cube” sensor
Coordinate finding magnet
Use “Hall cube” sensor developed at PSI. - 6 Hall sensors to measure 3 directions (2 sensors for each direction) - Sensors are mounted in a small space (200 ). - Cuboid structure makes the sensor orthogonal to each other.
Use a coordinate finding magnet to measure the sensor directions. It consists of Nd magnets which provide uniform field with known direction.

## Slide 8: Hall sensor
B
V
Hall voltage
e-
Hall effect
B
B//
Planar Hall Effect (∝ )
B⊥
Hall sensor - Simple, easy to use - Careful calibration is required (small non-linearity exists)
- Gain drift (by temperature etc.)  monitoring
- Planar Hall Effect (<0.2%)  cancel out by having two sensors in each direction with 90 deg. rotated.
Instability of one of our sensor was relatively large (<0.1%)
We broke one out of 6 sensor by accident. Fortunately, PHE for that sensor (Z) is expected to be negligible.
Minor issues with our sensor:

## Slide 9: Measurement system
Hallcube sensor
R arm
Φ arm
Z stage
Two moving arms + encoders Supersonic motors (works in B-field) + 3D printed arms (reinforced w/ CFRP)
Z stage + encoder ~3m long, toothed belt Granite table (stable) + CFRP tube (light)
Sensor Hallcube + current source (Keythley 6221) + Digital voltmeter (Agilent)
Control system EPICS (multiple devices, real time)

## Slide 10: Survey of the mapping machine
Colored circles: R arm rotation
Black circle: Phi arm rotation (center of R circles)
Position and tilt of the arms are measured in laser survey.
laser marker
Z position is found to vary by ~0.5mm due to tilt, but we can correct this based on the tilt parameters obtained in this survey.

## Slide 11: Survey of the mapping machine
Linearity of the stage is also measured and corrected.
Deviation in XY was ~1.5 mm at maximum.
After the corrections of tilt and Z linearity, deviation (RMS) of surveyed X,Y,Z positions from calculated position were ~200 in X,Y, ~30 in Z.
laser marker

## Slide 12: Sensor direction measurement
Nd
x
z
y
θ
Φ
sensor
We measured the B-field with no rotation and 90deg rotated.
no rotation
90deg rotated
Measured field correspond to BcosθcosΦ and BcosθsinΦ.  θ and Φ are calculated.
Stage is aligned
horizontal and parallel to Z
Nd
Sensor direction was measured by using Nd magnets which provides uniform field with known direction.
Hall
sensor
B=0.185 T

## Slide 13: Data taking
We did a full mapping in 1 day (very quick!).
30deg step in , 2 cm (or 1 cm) step in R, ~5 mm step in Z
Measured data points (X,Y)
Example of sensor readout for sensor#2 (Z)
(∝Bz)
Φ=0
r=0, 2, …26 cm

## Slide 14: Stability check
Temperature vs. time
Gain vs. time
Gain of Hall sensor may vary due to temperature etc.
We monitored the gain and offset by scanning over Z every 10 min. at COBRA center.
Variation of offset was negligible. Gain change of ~0.1% will be corrected.
Example of monitoring data

## Slide 15: Next steps
Calibration (VGauss) of Hall sensor Using dipole magnet + NMR probe (reference)
Interpolate discrete points to make full 3D map Previous method: B-spline (2D spline) fit in each R-Z planes at different φ New method:
Fit with solutions to Maxwell equations for generic solenoid
etc.
Fitting software used in Mu2e is adopted for MEG. Misalignment of the sensors can be included in the fit parameters.
Dipole magnet

## Slide 16: Summary
High precision (~0.1%) is required for the measurement of the gradient B-field of MEG II COBRA magnet.
We successfully performed mapping:
  3D “Hallcube” developed at PSI + moving system were used.
  Moving stage was carefully surveyed with good precision (e.g. ΔZ~30 ).
  Direction of the sensor was measured by dedicated setup with permanent magnet.
  Stability (gain variation of <0.1%) of the sensor was monitored every ~10 min.
Calibration of the sensor and interpolation of the data points will be done to make a final B-field map for MEG II.