## Slide 1: MEG II実験 液体キセノンガンマ線検出器における 光センサーの較正
東大素セ
家城佳, 他MEG IIコラボレーション
JSPS
Core-to-Core Program

## Slide 2: decay
MEG II experiment searches for cLFV decay, .
Sensitivity goal: (10 times better than MEG)
BSM prediction : (e.g. SUSY-seesaw)
If is found
Discovery new physics!
History of LFV upper limits

## Slide 3: MEG II experiment
900l LXe detector
μ+
γ
e+
Cylindrical drift chamber
+ Timing counter
Key concepts:
- High rate continuous beam at PSI ( - High resolution detectors to distinguish from accidental BG
e
back-to-back same timing
e
or
BG (Accidental)
Signal
Finally, all detectors were installed in 2018!
(w/ partial readout electronics)

## Slide 4: LXe detector
1 m
2 inch
12 mm
900L liquid Xe (LXe) scintillator to detect energy, position and timing of
In MEG II, entrance face is replaced from 216 PMTs (2 inches) to 4092 MPPCs (12x12 mm2)
Light collection uniformity and granularity improved!  x2 energy and position resolution improvement expected

## Slide 5: Pre-engineering run 2018
data taking in Nov.-Dec. 2018
- 17.6 MeV from radiative capture of on target with CW proton beam (first time in MEG II) - 40~55 MeV from & with beam
Measure detector resolution
Measure BG
Only ¼ of readout electronics (1.2 GHz waveform digitizer) was available.
p beam + Li target
beam + CH2 target
or

## Slide 6: Calibration of MPPC and PMT
To achieve good resolution, characteristics of sensors must be understood well.
Nphotons =
Gain: charge of 1 p.e.
EQF: Excess charge factor = crosstalk + afterpulse of MPPC
PDE: Photon Detection Efficiency
LEDs
α (241Am) on wires
Calibration sources:
LEDs  Gain, EQF sources  PDE

## Slide 7: Gain and EQF (MPPC)
Example of charge distribution
Gain = (1p.e. charge) – (0p.e. charge)
0 p.e.
1 p.e.
EQF (= crosstalk + afterpulse)
=
Estimated from fraction of 0p.e. events assuming Poisson statistics. P(0p.e.) = exp(-charge_mean)
Gain and EQF are measured in weak LED data.

## Slide 8: EQF and production lot
crosstalk
afterpulse
crosstalk, afterpulse vs. serial number
example of EQF vs. integration time
S. Kobayashi
fit with
CCT + CAP(1-exp(t-t0))
crosstalk
afterpulse
lot A
B
C
D
Effect of crosstalk and afterpulse can be measured separately by changing charge integration time.
Afterpulse was measured to be largely dependent on production lot, but it is not a problem as far as we measure it.
Vover=7V
~640 MPPCs

## Slide 9: Test of all MPPCs
Gain, EQF of all MPPCs were measured with LED after pre-eng. run.
High dark current channel: 19 channels Reason is unknown
Dead channel:
9 channels (0.2% of total) Most likely due to bad cables
Gain vs. serial no.
lot A
B
C
D
Outliers in gain distribution are mainly due to bad readout electronics channels, but some of them are not due to electronics:
Waveform of high dark rate ch.
These bad channels does not make a big effect on resolution.
Vover=7V
~4000 MPPCs

## Slide 10: Gain (PMT)
Gain vs. channel ID
Charge variance vs. charge mean
For PMTs, gain is derived from variance and mean of LED data.
Mean = gain * Np.e. * e Variance = gain2 * * e2 +
Np.e.
noise, LED variation
gain = slope of variance vs. mean
Bias voltage is adjusted to equalize the gain.
K. Toyoda

## Slide 11: PDE (MPPC), QE (PMT)
S. Kobayashi
PDE vs. incident angle
Measured this time
Measured in prototype test
expected curve
(Fresnel reflection)
PDE, QE was measured with 241Am source inside LXe detector.
PDE = Np.e.(measured) / Np.e.(expected in MC)
PDE (MPPC) vs. serial number
lot
A
B
C
D
QE (PMT) vs. channel
MPPC PDE was measured to be small (~7%), for the reason not understood yet.
It is measured to be small at large incident angle, as we observed in prototype test.
Effect on performance is expected to be small at least for energy and timing resolution.

## Slide 12: Summary
Calibration constants of MPPCs and PMTs are measured successfully. - Afterpulse of MPPC was found to differ largely in different production lot. - For MPPCs, dead channels (x9) and high dark rate channels (x19) were found. - PDE was measured to be small (~7%), compared to the values measured before construction (~20%). Effect on energy and timing resolution is expected to be small, and effect on position resolution is being evaluated.
Variation of calibration constants over time is an other important topic  next talk
Effect on resolution is expected to be minor.

## Slide 13: Crosscheck with current
gain*CTAP*PDE vs. current
gain*EQF*PDE
current*scale_factor
gain*EQF*PDE vs. serial number
gain*EQF*PDE

## Slide 14: Backup slides

## Slide 15: Gain monitoring
K. Toyoda